[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

Remington & Vernick Engineers
Project Name: Maclearie Park
Project # N/A
Order ID # Q3495
Test Name: Pesticide-TCL
A. Number of Samples and Date of Receipt:

2 Solid samples were received on 10/29/2025.

2 Water samples were received on 10/29/2025.
B. Parameters


According to the Chain of Custody document, the following analyses were requested:

Pesticide-TCL. This data package contains results for Pesticide-TCL.

C. Analytical Techniques:

The analysis was performed on instrument ECD_D.  The front column is ZB-MR1 which is 30 meters, 0.32 mm ID, 0. 5 um df,: Catalog # 7HM-G016-17. The rear column is ZB-MR2 which is 30 meters, 0.32 mm ID, 0.25 um df, Catalog #: 7HMG017- 11.The analysis was performed on instrument ECD_L.  The front column is ZB-MR1 which is 30 meters, 0.32 mm ID, 0. 5 um df,: Catalog # 7HM-G016-17. The rear column is ZB-MR2 which is 30 meters, 0.32 mm ID, 0.25 um df, Catalog #: 7HMG017- 11.The analysis of Pesticide-TCLs was based on method 8081B and extraction was done based on method 3510.
D. QA/ QC Samples:

The Holding Times were met for all analysis.
The Surrogate recoveries were met for all analysis.
The Retention Times were met for all analysis.
The MS recoveries met the requirements for all compounds.
The MSD recoveries met the requirements for all compounds.
The RPD were met for all analysis.
The Blank Spike met requirements for all compounds.

The Blank analysis did not indicate the presence of lab contamination.
The Initial Calibration met the requirements.

The Continuous Calibration File ID PD090949.D met the requirements except for Heptachlor is failing in 1st column, failing high but no positive hits in associated samples therefore no corrective action was taken. AND for Decachlorobiphenyl is failing in 2nd column, failing high but no positive hits in associated samples therefor no corrective action was taken.

The Continuous Calibration File ID PD090957.D met the requirements except for Heptachlor is failing in 1st column, failing high but no positive hits in associated samples therefore no corrective action was taken.
E. Additional Comments:
The soil samples  results  are based on a dry weight basis.
F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


