[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

Remington & Vernick Engineers
Project Name: Maclearie Park
Project # N/A
Order ID # Q3495
Test Name: VOC-TCLVOA-10
A. Number of Samples and Date of Receipt:

1 Solid sample was received on 10/29/2025.

2 Water samples were received on 10/29/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

VOC-TCLVOA-10. This data package contains results for VOC-TCLVOA-10.

C. Analytical Techniques:

The analysis performed on instrument MSVOA_X were done using GC column DB-624UI 20m 0.18mm 1.0 um. Cat#121-1324UI. The analysis performed on instrument MSVOA_Y were done using GC column Rxi-624SIL MS 30m, 0.25mm, 1.4 um, Cat. #13868.The analysis of VOC-TCLVOA-10 was based on method 8260-Low,8260D.
D. QA/ QC Samples:

The Holding Times were met for all analysis.
The Surrogate recoveries were met for all analysis except for TW-1025-1RE [Toluene-d8 - 77%] sample was reanalyzed to confirm the failure and reported.
The Internal Standards Areas were met for all analysis.
The Retention Times were met for all analysis.
The RPD for {VX1031WBSD01} with File ID: VX048420.D met criteria except for Tetrachloroethene[16%].
The Blank Spike for {VX1031WBS01} with File ID: VX048419.D met requirements for all compounds except for 1,3-Dichlorobenzene[81%], 1,4-Dichlorobenzene[81%] is failing low as corrective action associate sample reanalyzed under passing BS.

The Blank Spike Duplicate for {VY1030SBSD01} with File ID: VY023623.D met requirements for all compounds except for 1,2-Dichloropropane[123%] is failing high but no positive hit in associate sample therefore no corrective action taken.
The Blank analysis did not indicate the presence of lab contamination.
The Initial Calibration met the requirements.

The Continuous Calibration File ID VX048445.D met the requirements except for Dibromofluoromethane which is not our target compound, therefore no corrective action taken.

The Continuous Calibration File ID VY023620.D met the requirements except for Methyl Acetate and Styrene are failing high but no positive hit in associate sample therefore no corrective action taken.
The Tuning criteria met requirements.          

E. Additional Comments:
Samples for MS/MSD for VOC analysis were not provided with this set of samples. The Blank Spike Duplicate is reported with the data.
F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


