[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

Remington & Vernick Engineers
Project Name: Edison Landfill
Project # N/A
Order ID # Q3534
Test Name: SVOC-SIMGroup1
A. Number of Samples and Date of Receipt:

9 Water samples were received on 11/04/2025.
B. Parameters


According to the Chain of Custody document, the following analyses were requested:

SVOC-SIMGroup1. This data package contains results for SVOC-SIMGroup1.

C. Analytical Techniques:

The samples were analyzed on instrument BNA_N using GC Column ZB-SemiVolatiles Guardian which is 30 meters, 0.25 mm ID, 0.5 um df, Catalog # 7HG-G027-17-GGA. The analysis of SVOC-SIMGroup1 was based on method 8270-Modified and extraction was done based on method 3510.
D. QA/ QC Samples:

The Holding Times were met for all analysis.
The Surrogate recoveries were met for all analysis except for, 

MW-6 [Terphenyl-d14 - 131%], 

MW-11 [Terphenyl-d14 - 138%], 

MW-5 [Terphenyl-d14 - 135%], 

MW-EB-1 [Terphenyl-d14 - 136%], 

MW-EB-3 [Terphenyl-d14 - 135%],  

EB-1 [Terphenyl-d14 - 142%].
MW-6DL [Nitrobenzene-d5 - 132%],

MW-1DL [Terphenyl-d14 - 173%],

MW-11DL [Terphenyl-d14 - 134%], 

MW-EB-1DL [Terphenyl-d14 - 138%], This compound did not meet the NJDKQP criteria but met the in-house criteria.
And,

MW-6DL [2-Fluorobiphenyl - 165%, Terphenyl-d14 - 210%], 
MW-1DL [2-Fluorobiphenyl - 233%, Nitrobenzene-d5 - 165%], This compound did not meet the NJDKQP criteria and in-house criteria. Due to high concentration of compounds, this sample required dilution. Therefore, sample was reanalyzed with dilution and reported
.

The Internal Standards Areas were met for all analysis except for,
MW-6, MW-EB-1, Due to high concentration of compounds, this sample required dilution. Therefore, sample was reanalyzed with dilution and reported and MW-EB-1DL failed internal standard is not associated with reporting list, Therefor no further corrective action was taken.
The Retention Times were met for all analysis.
The RPD were met for all analysis.
The Blank Spike met requirements for all compounds.

The Blank Spike Duplicate met requirements for all compounds.
The Blank analysis did not indicate the presence of lab contamination.
The initial calibration met the requirements.
The Continuous Calibration File ID BN038174.D met the requirements except for 2,4,6-Tribromophenol. Failed surrogate is not associated with reporting list, Therefor no further corrective action was taken.
The Tuning criteria met requirements.          

Samples MW-6, MW-1, MW-11 and MW-EB-1 were diluted due to high concentrations.
E. Additional Comments:
The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.
F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


