ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS SEMI-VOLATILE ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY
ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS SEMI-VOLATILE ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY (CONTINUED)
	
	
	NA
	NO
	YES



	ORDER ID: Q3534

	MATRIX: Water

	METHOD: 8270E/3510

	


	
	
	NA
	NO
	YES

	1.
	Chromatograms Labeled/Compounds Identified.  (Field samples and Method Blanks)


	
	
	(

	2.
	GC/MS Tuning Specifications. DFTPP Meet Criteria.

(NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, 

CLP AND NJ)


	
	
	(

	3.
	GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series.


	
	
	(

	4.
	GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series.


	
	
	(

	5.
	GC/MS Calibration Requirements.

	
	(
	

	
	The %RSD is greater than 20% in the Method 8270-BF110525.M for Hexachlorocyclopentadiene is passing on Quadratic regression.

The %RSD is greater than 20% in the Method 8270-BG102421.M for 2-Nitrophenol 2-Nitroaniline, 2,6-Dinitrotoluene,   2,4-Dinitrophenol, 2,4-Dinitrotoluene, 4,6-Dinitro-2-methylphenolthese  Compounds are passing on Linear regression.

The %RSD is greater than 20% in the Initial Calibration (Method 8270-BP102925.M) for 26) 2-Nitrophenol, 2-Nitroaniline, 2,6-Dinitrotoluene, 3-Nitroaniline,2,4-Dinitrotoluene,Butylbenzylphthalate,Bis(2-ethylhexyl)phthalate. These  Compounds are passing on Linear regression 2,4-Dinitrophenol,4,6-Dinitro-2-methylphthalate are passing on Quadratic regression.

The Continuous Calibration File ID BG064699.D met the requirements except for 2,3,4,6-Tetrachlorophenol,2,4,6-Trichlorophenol,2,4-Dinitrotoluene,2,6-Dinitrotoluene,2-Nitrophenol,4,6-Dinitro-2-methylphenol,Hexachlorobutadiene,2,4,6-Tribromophenol and Nitrobenzene-d5, Associated samples does not have hit for these compounds, Therefor no further corrective  action was taken.

The Continuous Calibration File ID BP026076.D met the requirements except for 2,3,4,6-Tetrachlorophenol,2,4-Dinitrophenol,2,4-Dinitrotoluene,2,6-Dinitrotoluene,2-Nitroaniline,2-Nitrophenol,3-Nitroaniline,4,6-Dinitro-2-methylphenol, ,Butylbenzylphthalate,Di-n-octyl phthalate,Hexachlorocyclopentadiene,Pentachlorophenol and 2,4,6-Tribromophenol. Associated samples does not have hit for these compounds, Therefor no further corrective action was taken,


	
	
	

	6.
	Blank Contamination - If yes, list compounds and concentrations in each blank:


	
	(
	

	7.
	Surrogate Recoveries Meet Criteria


If not met, list those compounds and their recoveries which fall outside the acceptable ranges.

The Surrogate recoveries were met for all analysis except for PB170424BL [2,4,6-Tribromophenol - 113%], MW-6 [2,4,6-Tribromophenol - 126%], MW-11 [2,4,6-Tribromophenol - 125%], MW-5 [2,4,6-Tribromophenol - 116%], MW-EB-3 [2,4 and6-Tribromophenol - 113%]. This compound did not meet the NJDKQP criteria but met the in-house criteria.

	
	(
	

	8.
	Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria

If not met, list those compounds and their recoveries which fall outside the acceptable range.

The Blank Spike for {PB170424BS} with File ID: BP026078.D met requirements for all compounds except for 4-Chloroaniline[60%]. This compound did not meet the NJDKQP criteria but met the in-house criteria.

The Blank Spike Duplicate for {PB170424BSD} with File ID: BP026079.D met requirements for all compounds except for 4-Chloroaniline[59%]. This compound did not meet the NJDKQP criteria but met the in-house criteria.

	
	(
	

	9.
	Internal Standard Area/Retention Time Shift Meet Criteria

Comments: The Internal Standards Areas were met for all analysis.

	
	(
	

	10.
	Extraction Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(

	11.
	Analysis Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(


ADDITIONAL COMMENTS: 
The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.
_______________________________________



__________________________

QA REVIEW







Date

