[image: image1.png]AII|anoe

ECHNICA






284 Sheffield Street, Mountainside, NJ 7092, Phone: 908 789 8900, Fax: 908 789 8922


CASE NARRATIVE

Lockwood, Kessler & Bartlett, Inc.
Project Name: Syosset Landfill 2025
Project # N/A
Order ID # Q3560
Test Name: VOCMS Group1,Dissolved Metals Group5,Metals Group4,Alkalinity, Ammonia,Anions Group1,BOD5, Cyanide,pH,TDS,TKN,TOC,Total  Nitrogen,TSS, Turbidity
A. Number of Samples and Date of Receipt:

10 Water samples were received on 11/06/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

VOCMS Group1,Dissolved Metals Group5,Metals Group4,Alkalinity,Ammonia,Anions Group1,BOD5,Cyanide,pH,TDS,TKN,TOC,Total  Nitrogen,TSS,Turbidity. This data package contains results for VOCMS Group1(8260-Low),Dissolved Metals Group5 (6010D),Metals Group4(6010D),Alkalinity(SM2320 B),Ammonia(SM4500-NH3), Anions Group1(300.0),BOD5(SM5210 B),Cyanide(9012B),pH(9040C),TDS(SM2540 C),TKN(SM4500 N Org B or C),TOC(SM5310B),Total  Nitrogen(Cal),TSS(SM2540 D),Turbidity(SM2130 B).

C. Analytical Techniques:

VOCMS Group1 : The analysis performed on instrument MSVOA_N were done using GC column Rxi-624SIL MS 30m, 0.25mm, 1.4 um, Cat. #13868.The analysis of VOCMS Group1 was based on method 8260-Low.

Dissolved Metals Group5,Metals Group4 : The analysis of Dissolved Metals Group5,Metals Group4 was based on method 6010D and digestion based on method 3010 (waters).

Wetchem : The analysis of Anions Group1 was based on method 300.0, The analysis of Cyanide was based on method 9012B, The analysis of pH was based on method 9040C, The analysis of Total  Nitrogen was based on method Cal, The analysis of Turbidity was based on method SM2130 B, The analysis of Alkalinity was based on method SM2320 B, The analysis of TDS was based on method SM2540 C, The analysis of TSS was based on method SM2540 D, The analysis of TKN was based on method SM4500 N Org B or C, The analysis of Ammonia was based on method SM4500-NH3, The analysis of BOD5 was based on method SM5210 B and The analysis of TOC was based on method SM5310B.

D. QA/ QC Samples:

The Holding Times were met for all analysis except following Wetchem : SY-10D of pH, for SY-10I of pH, for SY-10S of pH, for SY-12D of pH and for SY-12I of pH as sample were receive out of holding time. SY-10IDL of Nitrate as samples Dilutions analyzed out of holding time.
The Surrogate recoveries were met for all analysis.
The Internal Standards Areas met the acceptable requirements.
The Retention Times were acceptable for all samples.
The MS recoveries met the requirements for all compounds except following Wetchem : The Matrix Spike (SY-10SMS) analysis met criteria for all compounds except for TOC due to sample matrix interference. The Matrix Spike (MANHOLEMS) analysis met criteria for all compounds except for Chloride due to sample matrix interference.
The MSD recoveries met the requirements for all compounds except following 

Dissolved Metals Group5,Metals Group4 : The Matrix Spike Duplicate (WATER-TREATMENT DISCHARGEMSD) analysis met criteria for all compounds except for Chromium and Iron Due to Chemical Interference during Digestion process.

 Wetchem : The Matrix Spike Duplicate(MANHOLEMSD) analysis met criteria for all compounds except for Chloride due to sample matrix interference.

The RPD recoveries met criteria.
The Blank Spike met requirements for all compounds.  
The Blank Spike Duplicate met requirements for all compounds.
The Blank analysis did not indicate the presence of lab contamination.
The Initial Calibration met the requirements.

The Continuous Calibration met the requirements except following 

VOCMS Group1 : The Continuous Calibration File ID VN088222.D met the requirements except for 2-Hexanone is failing high but no positive hit in associate sample therefore no corrective action taken.

The Tuning criteria met requirements.          

Wetchem : Sample SY-10D was diluted due to high concentrations for Ammonia as N, Chloride & Sample SY-10S was diluted due to high concentrations for Chloride & Sample SY-10I was diluted due to high concentrations for Chloride, Nitrate & Sample SY-12I was diluted due to high concentrations for Ammonia as N, TKN, Chloride, Sulfate & Sample SY-12IDL was diluted due to high concentrations for Ammonia as N, Chloride & Sample SY-12D was diluted due to high concentrations for Chloride, Nitrate, Sulfate & Sample SY-12DDL was diluted due to high concentrations for Chloride, Sulfate.

The Duplicate analysis met criteria for all compounds.
The Serial Dilution met the acceptable requirements.
E. Additional Comments:
Dissolved Metals Group5,Metals Group4 : Sample Q3560-01, Q3560-03, Q3560-05, Q3560-07 and Q3560-09 were analyzed as Total Metal and sample Q3560-02, Q3560-04, Q3560-06, Q3560-08 and Q3560-10 were analyzed as Dissolved Metal.

VOCMS Group1 : Samples for MS/MSD for VOC analysis were not provided with this set of samples. The Blank Spike Duplicate is reported with the data.

Trip Blank was not provided with this set of samples.

Alliance has analyzed samples for VOCMS Group1 by Method 8260D for Project “Syosset Landfill 2025”. Alliance  is not certified for trans-1,4-dichloro-2-butene compound with NJDEP for 8260D method. for VOCMS Group1 at the time when samples for Project “Syosset Landfill 2025 “were analyzed.

F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.
[image: image1.png]
I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


