[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

Remington & Vernick Engineers
Project Name: Edison Landfill
Project # N/A
Order ID # Q3569
Test Name: SVOC-TCL BNA -20
A. Number of Samples and Date of Receipt:

2 Water samples were received on 11/06/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

SVOC-TCL BNA -20. This data package contains results for SVOC-TCL BNA -20.

C. Analytical Techniques:

The samples were analyzed on instrument BNA_F using GC Column DB-UI 8270D which is 20 meters, 0.18 mm ID, 0.36 um df. The samples were analyzed on instrument BNA_P using GC Column ZB-SemiVolatiles Guardian which is 30 meters, 0.25 mm ID, 0.5 um df, Catalog # 7HG-G027-17-GGA. The analysis of SVOC-TCL BNA -20 was based on method 8270E and extraction was done based on method 3510.
D. QA/ QC Samples:

The Holding Times were met for all analysis.
The Surrogate recoveries were met for all analysis except for, 
MW-2 [2,4 and6-Tribromophenol - 122%]. This compound did not meet the NJDKQP criteria but met the in-house criteria.
The Internal Standards Areas were met for all analysis.
The Retention Times were met for all analysis.
The RPD were met for all analysis.
The Blank Spike for {PB170473BS} with File ID: BF144245.D met requirements for all compounds except for 3,3-Dichlorobenzidine[45%], 3-Nitroaniline[52%] and 4-Chloroaniline[33%]. This compound did not meet the NJDKQP criteria but met the in-house criteria.

The Blank Spike Duplicate for {PB170473BSD} with File ID: BF144246.D met requirements for all compounds except for 3,3-Dichlorobenzidine[45%], 3-Nitroaniline[50%] and 4-Chloroaniline[33%]. This compound did not meet the NJDKQP criteria but met the in-house criteria.
The Blank analysis did not indicate the presence of lab contamination.
The %RSD is greater than 20% in the Initial Calibration (Method 8270-BF110525.M)

for Hexachlorocyclopentadiene is passing on Quadratic regression.

The %RSD is greater than 20% in the Initial Calibration (Method 8270-BP102925.M) for 26) 2-Nitrophenol, 2-Nitroaniline, 2,6-Dinitrotoluene, 3-Nitroaniline,2,4-Dinitrotoluene,Butylbenzylphthalate,Bis(2-ethylhexyl)phthalate, These  Compounds are passing on Linear regression and 2,4-Dinitrophenol,4,6-Dinitro-2-methylphthalate are passing on Quadratic regression.

The Continuous Calibration File ID BF144243.D met the requirements except for 2,4-Dinitrophenol,Benzaldehyde and Hexachlorocyclopentadiene. Associated sample does not have hit for these compounds, Therefor no further corrective action was taken.

The Continuous Calibration File ID BP026090.D met the requirements except for 2,3,4,6-Tetrachlorophenol,2,4-Dinitrophenol,2,4-Dinitrotoluene,2-Nitrophenol,4,6-Dinitro-2-methylphenol,Butylbenzylphthalate,Di-n-octyl phthalate,Hexachlorocyclopentadiene and 2,4,6-Tribromophenol.  Associated sample does not have hit for these compounds, Therefor no further corrective action was taken.

The Continuous Calibration File ID BP026108.D met the requirements except for 2,3,4,6-Tetrachlorophenol,2,4-Dinitrophenol,2-Nitrophenol,4,6-Dinitro-2-methylphenol,Di-n-octyl phthalate and 2,4,6-Tribromophenol. Associated sample does not have hit for these compounds, Therefor no further corrective action was taken.
The Tuning criteria met requirements.          

Sample MW-12 was diluted due to high concentration.
E. Additional Comments:
The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.
F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


