ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS SEMI-VOLATILE ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY
ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS SEMI-VOLATILE ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY (CONTINUED)
	
	
	NA
	NO
	YES



	ORDER ID: Q3569

	MATRIX: Water

	METHOD: 8270E/3510

	


	
	
	NA
	NO
	YES

	1.
	Chromatograms Labeled/Compounds Identified.  (Field samples and Method Blanks)


	
	
	(

	2.
	GC/MS Tuning Specifications. DFTPP Meet Criteria.

(NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, 

CLP AND NJ)


	
	
	(

	3.
	GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series.


	
	
	(

	4.
	GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series.


	
	
	(

	5.
	GC/MS Calibration Requirements.

	
	(
	

	
	The %RSD is greater than 20% in the (Method 8270-BF110525.M) for Hexachlorocyclopentadiene is passing on Quadratic regression.

The %RSD is greater than 20% in the Initial Calibration (Method 8270-BP102925.M) for 26) 2-Nitrophenol, 2-Nitroaniline, 2,6-Dinitrotoluene, 3-Nitroaniline,2,4-Dinitrotoluene,Butylbenzylphthalate,Bis(2-ethylhexyl)phthalate, These  Compounds are passing on Linear regression and 2,4-Dinitrophenol,4,6-Dinitro-2-methylphthalate are passing on Quadratic regression.

The Continuous Calibration File ID BF144243.D met the requirements except for 2,4-Dinitrophenol,Benzaldehyde and Hexachlorocyclopentadiene. Associated sample does not have hit for these compounds, Therefor no further corrective action was taken.

The Continuous Calibration File ID BP026090.D met the requirements except for 2,3,4,6-Tetrachlorophenol,2,4-Dinitrophenol,2,4-Dinitrotoluene,2-Nitrophenol,4,6-Dinitro-2-methylphenol,Bis(2-ethylhexyl)phthalate,Butylbenzylphthalate,Di-n-octyl phthalate,Hexachlorocyclopentadiene and 2,4,6-Tribromophenol.  Associated sample does not have hit for these compounds, Therefor no further corrective action was taken.

The Continuous Calibration File ID BP026108.D met the requirements except for 2,3,4,6-Tetrachlorophenol,2,4-Dinitrophenol,2-Nitrophenol,4,6-Dinitro-2-methylphenol,Di-n-octyl phthalate and 2,4,6-Tribromophenol. Associated sample does not have hit for these compounds, Therefor no further corrective action was taken.


	
	
	

	6.
	Blank Contamination - If yes, list compounds and concentrations in each blank:


	
	(
	

	7.
	Surrogate Recoveries Meet Criteria


If not met, list those compounds and their recoveries which fall outside the acceptable ranges.

The Surrogate recoveries were met for all analysis except for MW-2 [2,4 and6-Tribromophenol - 122%]. This compound did not meet the NJDKQP criteria but met the in-house criteria.

	
	(
	

	8.
	Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria

If not met, list those compounds and their recoveries which fall outside the acceptable range.

The Blank Spike for {PB170473BS} with File ID: BF144245.D met requirements for all compounds except for 3,3-Dichlorobenzidine[45%], 3-Nitroaniline[52%] and 4-Chloroaniline[33%]. This compound did not meet the NJDKQP criteria but met the in-house criteria.

The Blank Spike Duplicate for {PB170473BSD} with File ID: BF144246.D met requirements for all compounds except for 3,3-Dichlorobenzidine[45%], 3-Nitroaniline[50%] and 4-Chloroaniline[33%]. This compound did not meet the NJDKQP criteria but met the in-house criteria.

	
	(
	

	9.
	Internal Standard Area/Retention Time Shift Meet Criteria

Comments: 


	
	
	(

	10.
	Extraction Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(

	11.
	Analysis Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(


ADDITIONAL COMMENTS: 
Sample MW-12 was diluted due to high concentration.

The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.
_______________________________________



__________________________

QA REVIEW







Date

