[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

Pacific Commercial Services Inc.
Project Name: Kilo Pier
Project # N/A
Order ID # Q3574
Test Name: SVOC-PAH
A. Number of Samples and Date of Receipt:

3 Water samples were received on 11/07/2025.
B. Parameters


According to the Chain of Custody document, the following analyses were requested:

SVOC-PAH. This data package contains results for SVOC-PAH.

C. Analytical Techniques:

The samples were analyzed on instrument BNA_F using GC Column DB-UI 8270D which is 20 meters, 0.18 mm ID, 0.36 um df. The analysis of SVOC-PAH was based on method 8270E and extraction was done based on method 3510.
D. QA/ QC Samples:

The Holding Times were met for all analysis.
The Surrogate recoveries were met for all analysis except for 304641-01 Liquid [Terphenyl-d14 - 36%], as per method one base surrogate allow to fail therefore no corrective action taken.
The Internal Standards Areas were met for all analysis.
The Retention Times were met for all analysis.
The RPD were met for all analysis.
The Blank Spike met requirements for all compounds.

The Blank Spike Duplicate met requirements for all compounds.
The Blank analysis did not indicate the presence of lab contamination.
The Initial Calibration met the requirements.

The Continuous Calibration met the requirements.
The Tuning criteria met requirements.  
Samples 304641-01 Liquid was diluted due oily matrix was analyzed with 20X dilution.
E. Additional Comments:
The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


