          [image: image2.png]AII|anoe

ECHNICA




         284 Sheffield Street, Mountainside, NJ 07092 Phone: 908 789 8900 Fax: 908 789 8922
GC ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY
          [image: image1.png]AII|anoe

ECHNICA




         284 Sheffield Street, Mountainside, NJ 07092 Phone: 908 789 8900 Fax: 908 789 8922
GC ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY (CONTINUED)
	
	
	NA
	NO
	YES



	ORDER ID: Q3586

	MATRIX: Solid

	METHOD: 8081B/3541

	


	
	
	NA
	NO
	YES

	1.
	Chromatograms Labeled/Compounds Identified.


	
	
	(

	2.
	Standard Summary Submitted.


	
	
	(

	3.
	Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis, 12 HOURS IF 8000 SERIES METHOD.

The Initial Calibration met the requirements.
The Continuous Calibration met the requirements.

 ASK  wdcalcomm " "  \* MERGEFORMAT 

	
	
	(

	 4.
	Blank Contamination - If yes, list compounds and concentrations in each blank:


	
	(
	

	 5.
	Surrogate Recoveries Meet Criteria


If not met, list those compounds and their recoveries which fall outside the acceptable ranges.

The Surrogate recoveries were met for all analysis except for SED-2 [Decachlorobiphenyl(1)25%, Decachlorobiphenyl(2)27%]these compounds did not meet the NJDKQP criteria but met the in-house criteria.


	
	(
	

	 6.
	Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria

If not met, list those compounds and their recoveries which fall outside the acceptable range.

The MS recoveries met the requirements for all compounds.

The MSD recoveries met the requirements for all compounds. 

The Blank Spike met requirements for all compounds.

The RPD were met for all analysis.

	
	(
	

	7.
	Retention Time Shift Meet Criteria (if applicable)

Comments: 


	
	
	(

	8.
	Extraction Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(

	9.
	Analysis Holding Time Met 






If not met, list those compounds and their recoveries which fall outside the acceptable range.


	
	
	(


ADDITIONAL COMMENTS: 
The soil samples  results  are based on a dry weight basis.
_______________________________________



__________________________

QA REVIEW







Date

[image: image1.png][image: image2.png]