[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

Tris Pharma, Inc.
Project Name: Quarterly
Project # N/A
Order ID # Q3676
Test Name: VOCMS Group1
A. Number of Samples and Date of Receipt:

4 Water samples were received on 11/19/2025.
B. Parameters


According to the Chain of Custody document, the following analyses were requested:

VOCMS Group1. This data package contains results for VOCMS Group1.

C. Analytical Techniques:

The analysis performed on instrument MSVOA_N were done using GC column Rxi-624SIL MS 30m, 0.25mm, 1.4 um, Cat. #13868.The analysis of VOCMS Group1 was based on method 624.1.
D. QA/ QC Samples:

The Holding Times were met for all analysis.
The Surrogate recoveries were met for all analysis.
The Internal Standards Areas were met for all analysis except for OUTFALL-DSN-002, Due to high concentration of compounds, this sample required dilution. Therefore, sample was reanalyzed with dilution and reported.
The Retention Times were met for all analysis.
The MS {Q3676-02MS} with File ID: VN088324.D recoveries met the requirements for all compounds except for Acetone[240%] due to matrix interference.
The MSD {Q3676-03MSD} with File ID: VN088325.D recoveries met the requirements for all compounds except for Acetone[230%] due to matrix interference.
The RPD for {Q3676-03MSD} with File ID: VN088325.D met criteria except for Ethyl Acetate[22%], Methylene Chloride[22%] and n-amyl Acetate[33%] due to difference in results of MS and MSD.
The Blank Spike met requirements for all compounds.

The Blank analysis did not indicate the presence of lab contamination.
The Initial Calibration met the requirements.

The Continuous Calibration met the requirements.
The Tuning criteria met requirements.          

Sample OUTFALL-DSN-002 was diluted due to high concentration.
E. Additional Comments:
Trip Blank was not provided with this set of samples.

Alliance has analyzed samples for VOCMS Group1 by Method 624.1  for Project “Quarterly”. Alliance  is not certified for Isopropyl Acetate and N-amyl acetate compounds with NJDEP for 624.1  method. for VOCMS Group1 at the time when samples for Project “Quarterly “were analyzed.
F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


