[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

Ardmore Chemical
Project Name: PVSC Monthly 2025
Project # N/A
Order ID # Q3700
Test Name: SVOCMS Group1
A. Number of Samples and Date of Receipt:

3 Water samples were received on 11/20/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

SVOCMS Group1. This data package contains results for SVOCMS Group1.

C. Analytical Techniques:

The samples were analyzed on instrument BNA_F using GC Column DB-UI 8270D which is 20 meters, 0.18 mm ID, 0.36 um df. The analysis of SVOCMS Group1 was based on method 625.1 and extraction was done based on method 3510.
D. QA/ QC Samples:

The Holding Times were met for all analysis.
The Surrogate recoveries were met for all analysis except for EFF-WW [2-Fluorophenol - 52%, Phenol-d6 - 36%], EFF-WWMS [2-Fluorophenol - 56%, Phenol-d6 - 41%], EFF-WWMSD [2-Fluorophenol - 58% and Phenol-d6 - 40%] failing low but confirm with MS/MSD due to matrix interference therefore no corrective action is needed.
The Internal Standards Areas were met for all analysis.
The Retention Times were met for all analysis.
The MS recoveries met the requirements for all compounds.
The MSD recoveries met the requirements for all compounds.
The RPD for {Q3700-06MSD} with File ID: BF144354.D met criteria except for Benzidine[29%], due to difference in results of MS and MSD.
The Blank Spike for {PB170730BS} with File ID: BF144350.D met requirements for all compounds except for Benzidine[9%]marginally failing low therefore no corrective action was taken.
The Blank analysis did not indicate the presence of lab contamination.
The Initial Calibration met the requirements.
The Continuous Calibration met the requirements..
The Tuning criteria met requirements.          

E. Additional Comments:
F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


