ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS SEMI-VOLATILE ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY
ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS SEMI-VOLATILE ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY (CONTINUED)
	
	
	NA
	NO
	YES



	ORDER ID: Q3700

	MATRIX: Water

	METHOD: 625.1/3510

	


	
	
	NA
	NO
	YES

	1.
	Chromatograms Labeled/Compounds Identified.  (Field samples and Method Blanks)


	
	
	(

	2.
	GC/MS Tuning Specifications. DFTPP Meet Criteria.

(NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, 

CLP AND NJ)


	
	
	(

	3.
	GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series.


	
	
	(

	4.
	GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series.


	
	
	(

	5.
	GC/MS Calibration Requirements.

	
	(
	

	
	The %RSD is greater than 20% in the Method 8270-BF110525.M for Hexachlorocyclopentadiene is passing on Quadratic regression.

The Continuous Calibration File ID BF144348.D met the requirements except for Benzidine, The associate samples have no positive hit for these compounds therefore no corrective action was taken. and Terphenyl-d14 marginally low therefore no corrective action was taken.


	
	
	

	6.
	Blank Contamination - If yes, list compounds and concentrations in each blank:


	
	(
	

	7.
	Surrogate Recoveries Meet Criteria


If not met, list those compounds and their recoveries which fall outside the acceptable ranges.

The Surrogate recoveries were met for all analysis except for EFF-WW [2-Fluorophenol - 52%, Phenol-d6 - 36%], EFF-WWMS [2-Fluorophenol - 56%, Phenol-d6 - 41%], EFF-WWMSD [2-Fluorophenol - 58% and Phenol-d6 - 40%] failing low but confirm with MS/MSD due to matrix interference therefore no corrective action is needed.

	
	(
	

	8.
	Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria

If not met, list those compounds and their recoveries which fall outside the acceptable range.

The Blank Spike for {PB170730BS} with File ID: BF144350.D met requirements for all compounds except for Benzidine[9%]marginally failing low therefore no corrective action was taken.The RPD for {Q3700-06MSD} with File ID: BF144354.D met criteria except for Benzidine[29%], due to difference in results of MS and MSD.

	
	(
	

	9.
	Internal Standard Area/Retention Time Shift Meet Criteria

Comments: 


	
	
	(

	10.
	Extraction Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(

	11.
	Analysis Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(


ADDITIONAL COMMENTS: 
As per method, MS/MSD is required to be performed with sample analysis. However,

Lab did not receive sufficient volume to perform the MS/MSD therefore MS/MSD were

not performed for this project. However, Lab has performed LCS/LCSD instead.
_______________________________________



__________________________

QA REVIEW







Date

