[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

Ardmore Chemical
Project Name: PVSC Monthly 2025
Project # N/A
Order ID # Q3700
Test Name: VOC-PP
A. Number of Samples and Date of Receipt:

3 Water samples were received on 11/20/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

VOC-PP. This data package contains results for VOC-PP.

C. Analytical Techniques:

The analysis performed on instrument MSVOA_N were done using GC column Rxi-624SIL MS 30m, 0.25mm, 1.4 um, Cat. #13868.The analysis of VOC-PP was based on method 624.1.
D. QA/ QC Samples:

The Holding Times were met for all analysis.
The Surrogate recoveries were met for all analysis.
The Internal Standards Areas were met for all analysis.
The Retention Times were met for all analysis.
The MS {Q3700-02MS} with File ID: VN088360.D recoveries met the requirements for all compounds except for 2-Chloroethyl vinyl ether[0%], Acrolein[164%] due to matrix interference.
The MSD recoveries met the requirements for all compounds.
The RPD for {Q3700-03MSD} with File ID: VN088367.D met criteria except for 2-Chloroethyl vinyl ether[200%] due to difference in results of MS and MSD.
The Blank Spike for {VN1125WBS01} with File ID: VN088354.D met requirements for all compounds except for Acrolein[160%] is failing high and associate sample having hit of Acrolein but below CRQL therefore no corrective action taken.

The Blank Spike for {VN1126WBS01} with File ID: VN088363.D met requirements for all compounds except for Acrolein[150%] is failing high and associate QC passing within limit therefore no corrective action taken.
The Blank analysis did not indicate the presence of lab contamination.
The Initial Calibration met the requirements.

The Continuous Calibration met the requirements.
The Tuning criteria met requirements.          

Samples EFF-WW was diluted due to foamy nature of sample.
E. Additional Comments:
Trip Blank was not provided with this set of samples.
F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


