[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

Roman E&G Corp
Project Name: AQUA Woolwich Pump Station 22-531
Project # N/A
Order ID # Q3725
Test Name: SVOCMS NJ CleanGroup New
A. Number of Samples and Date of Receipt:

1 Solid sample was received on 11/25/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

SVOCMS NJ CleanGroup New. This data package contains results for SVOCMS NJ CleanGroup New.

C. Analytical Techniques:

The samples were analyzed on instrument BNA_F using GC Column DB-UI 8270D which is 20 meters, 0.18 mm ID, 0.36 um df. The analysis of SVOCMS NJ CleanGroup New was based on method 8270E and extraction was done based on method 3541.
D. QA/ QC Samples:

The Holding Times were met for all analysis.
The Surrogate recoveries were met for all analysis.
The Internal Standards Areas were met for all analysis.
The Retention Times were met for all analysis.
The MS {Q3735-01MS} with File ID: BF144392.D recoveries met the requirements for all compounds except for 3,3-Dichlorobenzidine[63%] and 4,6-Dinitro-2-methylphenol[52%]. These compounds did not meet the NJDKQP criteria but met the in-house criteria due to matrix interference.
The MSD {Q3735-01MSD} with File ID: BF144393.D recoveries met the requirements for all compounds except for 3,3-Dichlorobenzidine[68%] and 4,6-Dinitro-2-methylphenol[43%]. These compounds did not meet the NJDKQP criteria but met the in-house criteria due to matrix interference.
The RPD for {Q3735-01MSD} with File ID: BF144393.D met criteria except for 2,4-Dinitrophenol[38%], Hexachlorocyclopentadiene[21%]. Due to difference in MS and MSD concentrations.
The Blank Spike for {PB170763BS} with File ID: BF144386.D met requirements for all compounds except for Benzidine[19%]. This compound did not meet the NJDKQP criteria but met the in-house criteria.

The Blank analysis did not indicate the presence of lab contamination.
The %RSD is greater than 20% in the Method 8270-BF110525.M for Hexachlorocyclopentadiene, this compound is passing on Quadratic regression.

The Continuous Calibration File ID BF144382.D met the requirements except for Benzaldehyde and Benzidine. But associated samples have no positive hit therefore no corrective action was taken.
The Tuning criteria met requirements.          

E. Additional Comments:
The soil samples  results  are based on a dry weight basis.

The time of sampling were not listed in the COC. 

The temperature of the samples at the time of receipt was 16.1°C.
F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


