ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS SEMI-VOLATILE ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY
ALLIANCE 284 Sheffield Street, Mountainside New Jersey 07092

NEW JERSEY LAB ID#: 20012: NEW YORK LAB ID#: 11376

GC/MS SEMI-VOLATILE ANALYSIS CONFORMANCE/NON-CONFORMANCE SUMMARY (CONTINUED)
	
	
	NA
	NO
	YES



	ORDER ID: Q3725

	MATRIX: Solid

	METHOD: 8270E/3541

	


	
	
	NA
	NO
	YES

	1.
	Chromatograms Labeled/Compounds Identified.  (Field samples and Method Blanks)


	
	
	(

	2.
	GC/MS Tuning Specifications. DFTPP Meet Criteria.

(NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, 

CLP AND NJ)


	
	
	(

	3.
	GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series.


	
	
	(

	4.
	GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series.


	
	
	(

	5.
	GC/MS Calibration Requirements.

	
	(
	

	
	The %RSD is greater than 20% in the Method 8270-BF110525.M for Hexachlorocyclopentadiene, this compound is passing on Quadratic regression.

The Continuous Calibration File ID BF144382.D met the requirements except for Benzaldehyde and Benzidine. But associated samples have no positive hit therefore no corrective action was taken.


	
	
	

	6.
	Blank Contamination - If yes, list compounds and concentrations in each blank:


	
	(
	

	7.
	Surrogate Recoveries Meet Criteria


If not met, list those compounds and their recoveries which fall outside the acceptable ranges.


	
	
	(

	8.
	Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria

If not met, list those compounds and their recoveries which fall outside the acceptable range.

The MS {Q3735-01MS} with File ID: BF144392.D recoveries met the requirements for all compounds except for 3,3-Dichlorobenzidine[63%] and 4,6-Dinitro-2-methylphenol[52%]. These compounds did not meet the NJDKQP criteria but met the in-house criteria due to matrix interference.

The MSD {Q3735-01MSD} with File ID: BF144393.D recoveries met the requirements for all compounds except for 3,3-Dichlorobenzidine[68%] and 4,6-Dinitro-2-methylphenol[43%]. These compounds did not meet the NJDKQP criteria but met the in-house criteria due to matrix interference.

The Blank Spike for {PB170763BS} with File ID: BF144386.D met requirements for all compounds except for Benzidine[19%]. This compound did not meet the NJDKQP criteria but met the in-house criteria.

The RPD for {Q3735-01MSD} with File ID: BF144393.D met criteria except for 2,4-Dinitrophenol[38%], Hexachlorocyclopentadiene[21%]. Due to difference in MS and MSD concentrations.

	
	(
	

	9.
	Internal Standard Area/Retention Time Shift Meet Criteria

Comments: 


	
	
	(

	10.
	Extraction Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(

	11.
	Analysis Holding Time Met 






If not met, list number of days exceeded for each sample:


	
	
	(


ADDITIONAL COMMENTS: 
The soil samples  results  are based on a dry weight basis.

The time of sampling were not listed in the COC. 

The temperature of the samples at the time of receipt was 16.1°C.
_______________________________________



__________________________

QA REVIEW







Date

