[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

Tetra Tech NUS, Inc.
Project Name: CTO WA194808 Dahlgren VA - DOD VELAP
Project # N/A
Order ID # Q3776
Test Name: SVOC-TCL BNA -20
A. Number of Samples and Date of Receipt:

1 Water sample was received on 12/04/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

SVOC-TCL BNA -20. This data package contains results for SVOC-TCL BNA -20.

C. Analytical Techniques:

The samples were analyzed on instrument BNA_P using GC Column ZB-Semi Volatiles Guardian which is 30 meters, 0.25 mm ID, 0.5 um df, Catalog # 7HG-G027-17-GGA. The analysis of SVOC-TCL BNA -20 was based on method 8270E and extraction was done based on method 3510.
D. QA/ QC Samples:

The Holding Times were met for all analysis.
The Surrogate recoveries were met for all analysis.
The Internal Standards Areas were met for all analysis.
The Retention Times were met for all analysis.
The RPD were met for all analysis.
The Blank Spike met requirements for all compounds.

The Blank Spike Duplicate met requirements for all compounds.
The Blank analysis did not indicate the presence of lab contamination.
The Initial Calibration met the requirements.

The Continuous Calibration File ID BP026250.D met the requirements except for Benzaldehyde, failing high but The associate samples have no positive hit for these compounds therefore no corrective action was taken.
The Tuning criteria met requirements.          

E. Additional Comments:
The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.
F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


