[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

JACOBS Engineering Group, Inc.
Project Name: Former Schlumberger STC PTC Site  D3868221
Project # N/A
Order ID # Q3787
Test Name: VOCMS Group3
A. Number of Samples and Date of Receipt:

11 Water samples were received on 12/04/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

VOCMS Group3. This data package contains results for VOCMS Group3.

C. Analytical Techniques:

The analysis performed on instrument MSVOA_X were done using GC column DB-624UI 20m 0.18mm 1.0 um. Cat#121-1324UI. The analysis of VOCMS Group3 was based on method 8260-Low.
D. QA/ QC Samples:

The Holding Times were met for all analysis.
The Surrogate recoveries were met for all analysis except for OWBR-02-170-120425 [4-Bromofluorobenzene - 186%, Dibromofluoromethane - 142%], OWBR-02-170-120425-FD [1,2-Dichloroethane-d4 - 142%, 4-Bromofluorobenzene - 171%, Dibromofluoromethane - 134%], OW-03B-51.5-120425-FD [1,2-Dichloroethane-d4 - 140%, 4-Bromofluorobenzene - 155%] and OW-08B-72.5-120425-FD [Dibromofluoromethane - 136%] these compounds did not meet the NJDKQP criteria but met the in-house criteria while OWBR-02-170-120425RE [4-Bromofluorobenzene - 125%], OWBR-02-170-120425-FDRE [4-Bromofluorobenzene - 122%], OW-03B-51.5-120425 [1,2-Dichloroethane-d4 - 128%], OW-03B-51.5-120425RE [4-Bromofluorobenzene - 122%], OW-03B-51.5-120425-FDRE [4-Bromofluorobenzene - 125%] and OW-08B-72.5-120425-FDRE [4-Bromofluorobenzene - 123%] these compounds did not meet the in-house criteria but met NJDKQP criteria, samples were reanalyzed to confirm the failure and reported.

The Internal Standards Areas were met for all analysis.
The Retention Times were met for all analysis.
The MS recoveries met the requirements for all compounds.
The MSD recoveries met the requirements for all compounds.
The RPD were met for all analysis.
The Blank Spike met requirements for all compounds.

The Blank analysis did not indicate the presence of lab contamination.
The Initial Calibration met the requirements.

The Continuous Calibration met the requirements.
The Tuning criteria met requirements.          

E. Additional Comments:
F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


