[image: image1.png]AII|anoe

ECHNICA



[image: image1.png]

CASE NARRATIVE

Tully Construction Co., Inc.
Project Name: MTA 26 Stations - Pierrepont
Project # N/A
Order ID # Q3790
Test Name: VOC-TCLVOA-10
A. Number of Samples and Date of Receipt:

1 Solid sample was received on 12/05/2025.

B. Parameters


According to the Chain of Custody document, the following analyses were requested:

VOC-TCLVOA-10. This data package contains results for VOC-TCLVOA-10.

C. Analytical Techniques:

The analysis performed on instrument MSVOA_W were done using GC column Rxi-624SIL MS 30m, 0.25mm, 1.4 um, Cat. #13868.The analysis of VOC-TCLVOA-10 was based on method 8260D.
D. QA/ QC Samples:

The Holding Times were met for all analysis.
The Surrogate recoveries were met for all analysis except for 304 FURMAN SOIL [Toluene-d8 - 65%]. As corrective action lab renalyzed Vial B for this sample but it was not purged, therefore Vial A is reported as final results. 
The Internal Standards Areas were met for all analysis except for 304 FURMAN SOIL. As corrective action lab renalyzed Vial B for this sample but it was not purged, therefore Vial A is reported as final results. 
The Retention Times were met for all analysis.
The RPD for {VW1208SBSD01} with File ID: VW032589.D met criteria except for Bromochloromethane[34%]. Due to difference in BS and BSD concentrations. 
The Blank Spike met requirements for all compounds.

The Blank Spike Duplicate for {VW1208SBSD01} with File ID: VW032589.D met requirements for all compounds except for Bromochloromethane[145%]. Filing high but associated samples have no positive hit for this compound, therefore no corrective action was taken.  
The Blank analysis did not indicate the presence of lab contamination.
The Initial Calibration met the requirements.

The Continuous Calibration met the requirements.
The Tuning criteria met requirements.          

E. Additional Comments:
As per special requirement for this project form-1 and Hit Summary are reported in mg/kg.

Samples for MS/MSD for VOC analysis were not provided with this set of samples. The Blank Spike Duplicate is reported with the data.
Trip Blank was not provided with this set of samples.

The soil samples  results  are based on a dry weight basis.
F. Manual Integration Comments:
Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________


