CASE NARRATIVE

### Arecon Ltd.

### Project Name: 40421

### Project # N/A

### Chemtech Project # J3137

Test Name: TO-15

A. Number of Samples and Date of Receipt:

1 Air sample was received on 05/24/2018.

B. Parameters

According to the Chain of Custody document, the following analyses were requested:

TO-15. This data package contains results for TO-15.

C. Analytical Techniques:

The analysis performed on instrument MSVOA_L were done using GC column RTX-1, which is 60 meters, 0.32 mm id, 1.0 um df, Restek Cat. #10157. The Trap was supplied by Entech, glass bead and Tenax , Entech 7100A Preconcentrator.The analysis of TO-15 was based on method TO-15.

D. QA/ QC Samples:

The Holding Times were met for all analysis.

The Surrogate recoveries met the acceptable criteria.

The Internal Standards Areas met the acceptable requirements.

The Retention Times were acceptable for all samples.

The RPD for {J3193-01DUP} with File ID: VL032033.D recoveries met criteria except for 2-Butanone[29.6%] .

The Blank Spike met requirements for all samples .

The Blank analysis indicated presence of Acetone[0.9ug/m3] Methylene Chloride[1.25 ug/m3] FileID:VL032026.D{VL0601ABL01} due to possible lab contamination.

The Initial Calibration met the requirements except for Dichlorodifluoromethane , Naphthalene had more than 30% RSD in the Initial Calibration with dated 05/31/2018 with L Instrument , as per method two compounds as allowed to be fail .Therefore no corrective action was required.

The ICV with file id VL031422.D had % DEV. more than 30 %( 30.5 %) for Dichlorodifluoromethane The associate samples have no positive hit for this compound less than CRQL; therefore no corrective action was required.   .

The Continuous Calibration met the requirements .

The Tuning criteria met requirements.

| Sample ID |  |  | File ID |  | Parameter |  | Review By |  | Review On |  | Supervised By |  | Supervised On |  | Reason |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| VSTDICCC010 |  |  | VL032016.D |  | Chlorobenzene-d5 |  | sam |  | 6/3/2018 2:13:17 PM |  | MMDadoda |  | 6/4/2018 10:54:52 AM |  | Peak Integrated by Software incorrectly |  |  |  |
| VSTDICCC010 |  |  | VL032016.D |  | Hexachloro-1,3-Butadiene |  | sam |  | 6/3/2018 2:13:17 PM |  | MMDadoda |  | 6/4/2018 10:54:52 AM |  | Peak Integrated by Software incorrectly |  |  |  |
| VSTDICCC010 |  |  | VL032016.D |  | m/p-Xylene |  | sam |  | 6/3/2018 2:13:17 PM |  | MMDadoda |  | 6/4/2018 10:54:52 AM |  | Peak Integrated by Software incorrectly |  |  |  |
| VSTDICC002 |  |  | VL032017.D |  | 1,1,1,2-Tetrachloroethane |  | sam |  | 6/3/2018 2:13:24 PM |  | MMDadoda |  | 6/4/2018 10:54:53 AM |  | Peak Integrated by Software incorrectly |  |  |  |
| VSTDICC002 |  |  | VL032017.D |  | 1,2,4-Trichlorobenzene |  | sam |  | 6/3/2018 2:13:24 PM |  | MMDadoda |  | 6/4/2018 10:54:53 AM |  | Peak Integrated by Software incorrectly |  |  |  |
| VSTDICC002 |  |  | VL032017.D |  | Benzyl Chloride |  | sam |  | 6/3/2018 2:13:24 PM |  | MMDadoda |  | 6/4/2018 10:54:53 AM |  | Peak Integrated by Software incorrectly |  |  |  |
| VSTDICC002 |  |  | VL032017.D |  | m/p-Xylene |  | sam |  | 6/3/2018 2:13:24 PM |  | MMDadoda |  | 6/4/2018 10:54:53 AM |  | Peak Integrated by Software incorrectly |  |  |  |
| VSTDICC002 |  |  | VL032017.D |  | Naphthalene |  | sam |  | 6/3/2018 2:13:24 PM |  | MMDadoda |  | 6/4/2018 10:54:53 AM |  | Peak Integrated by Software incorrectly |  |  |  |
| VSTDICC002 |  |  | VL032017.D |  | tert-Butyl alcohol |  | sam |  | 6/3/2018 2:13:24 PM |  | MMDadoda |  | 6/4/2018 10:54:53 AM |  | Peak Integrated by Software incorrectly |  |  |  |
| VSTDICC002 |  |  | VL032017.D |  | Tetrachloroethene |  | sam |  | 6/3/2018 2:13:24 PM |  | MMDadoda |  | 6/4/2018 10:54:53 AM |  | Peak Integrated by Software incorrectly |  |  |  |
| VSTDICC002 |  |  | VL032017.D |  | Tetrahydrofuran |  | sam |  | 6/3/2018 2:13:24 PM |  | MMDadoda |  | 6/4/2018 10:54:53 AM |  | Peak Integrated by Software incorrectly |  |  |  |
| VSTDICC0.5 |  | VL032019.D |  | Styrene |  | sam |  | 6/3/2018 2:13:31 PM |  | MMDadoda |  | 6/4/2018 10:54:56 AM |  | Peak Integrated by Software incorrectly |  |  |  |  |
| VSTDICC0.5 |  | VL032019.D |  | t-1,3-Dichloropropene |  | sam |  | 6/3/2018 2:13:31 PM |  | MMDadoda |  | 6/4/2018 10:54:56 AM |  | Peak Integrated by Software incorrectly |  |  |  |  |
| VSTDICC0.5 |  | VL032019.D |  | tert-Butyl alcohol |  | sam |  | 6/3/2018 2:13:31 PM |  | MMDadoda |  | 6/4/2018 10:54:56 AM |  | Peak Integrated by Software incorrectly |  |  |  |  |
| VSTDICC0.5 |  | VL032019.D |  | Tetrachloroethene |  | sam |  | 6/3/2018 2:13:31 PM |  | MMDadoda |  | 6/4/2018 10:54:56 AM |  | Peak Integrated by Software incorrectly |  |  |  |  |
| VSTDICC0.5 |  | VL032019.D |  | Tetrahydrofuran |  | sam |  | 6/3/2018 2:13:31 PM |  | MMDadoda |  | 6/4/2018 10:54:56 AM |  | Peak Integrated by Software incorrectly |  |  |  |  |
| VSTDICC0.5 |  | VL032019.D |  | Trichloroethene |  | sam |  | 6/3/2018 2:13:31 PM |  | MMDadoda |  | 6/4/2018 10:54:56 AM |  | Peak Integrated by Software incorrectly |  |  |  |  |
| VSTDICC0.1 |  | VL032020.D |  | 1,1,1-Trichloroethane |  | sam |  | 6/3/2018 2:13:36 PM |  | MMDadoda |  | 6/4/2018 10:54:57 AM |  | Peak Integrated by Software incorrectly |  |  |  |  |
| VSTDICC0.1 |  | VL032020.D |  | 1,1,2,2-Tetrachloroethane |  | sam |  | 6/3/2018 2:13:36 PM |  | MMDadoda |  | 6/4/2018 10:54:57 AM |  | Peak Integrated by Software incorrectly |  |  |  |  |
| VSTDICC0.1 |  | VL032020.D |  | Carbon Tetrachloride |  | sam |  | 6/3/2018 2:13:36 PM |  | MMDadoda |  | 6/4/2018 10:54:57 AM |  | Peak Integrated by Software incorrectly |  |  |  |  |
| VSTDICC0.1 |  | VL032020.D |  | Tetrachloroethene |  | sam |  | 6/3/2018 2:13:36 PM |  | MMDadoda |  | 6/4/2018 10:54:57 AM |  | Peak Integrated by Software incorrectly |  |  |  |  |
| VSTDICC0.1 |  | VL032020.D |  | Trichloroethene |  | sam |  | 6/3/2018 2:13:36 PM |  | MMDadoda |  | 6/4/2018 10:54:57 AM |  | Peak Integrated by Software incorrectly |  |  |  |  |
| VSTDICC0.1 |  | VL032020.D |  | Vinyl Chloride |  | sam |  | 6/3/2018 2:13:36 PM |  | MMDadoda |  | 6/4/2018 10:54:57 AM |  | Peak Integrated by Software incorrectly |  |  |  |  |
| VSTDICC0.03 |  | VL032021.D |  | 1,1,1-Trichloroethane |  | sam |  | 6/3/2018 2:14:21 PM |  | MMDadoda |  | 6/4/2018 10:54:58 AM |  | Peak Integrated by Software incorrectly |  |  |  |  |
| VSTDICC0.03 |  | VL032021.D |  | 1,1,2,2-Tetrachloroethane |  | sam |  | 6/3/2018 2:14:21 PM |  | MMDadoda |  | 6/4/2018 10:54:58 AM |  | Peak Integrated by Software incorrectly |  |  |  |  |
| VSTDICC0.03 |  | VL032021.D |  | Carbon Tetrachloride |  | sam |  | 6/3/2018 2:14:21 PM |  | MMDadoda |  | 6/4/2018 10:54:58 AM |  | Peak Integrated by Software incorrectly |  |  |  |  |
| VSTDICC0.03 |  | VL032021.D |  | Tetrachloroethene |  | sam |  | 6/3/2018 2:14:21 PM |  | MMDadoda |  | 6/4/2018 10:54:58 AM |  | Peak Integrated by Software incorrectly |  |  |  |  |
| VSTDICC0.03 |  | VL032021.D |  | Trichloroethene |  | sam |  | 6/3/2018 2:14:21 PM |  | MMDadoda |  | 6/4/2018 10:54:58 AM |  | Peak Integrated by Software incorrectly |  |  |  |  |
| VSTDICC0.03 |  | VL032021.D |  | Vinyl Chloride |  | sam |  | 6/3/2018 2:14:21 PM |  | MMDadoda |  | 6/4/2018 10:54:58 AM |  | Peak Integrated by Software incorrectly |  |  |  |  |
| VSTDICC015 |  | VL032022.D |  | m/p-Xylene |  | sam |  | 6/3/2018 2:13:42 PM |  | MMDadoda |  | 6/4/2018 10:55:09 AM |  | Peak Integrated by Software incorrectly |  |  |  |  |
| VSTDICV010 |  | VL032023.D |  | Chlorobenzene-d5 |  | sam |  | 6/4/2018 10:17:13 AM |  | MMDadoda |  | 6/4/2018 10:55:00 AM |  | Peak Integrated by Software incorrectly |  |  |  |  |
| VSTDICV010 |  | VL032023.D |  | m/p-Xylene |  | sam |  | 6/4/2018 10:17:13 AM |  | MMDadoda |  | 6/4/2018 10:55:00 AM |  | Peak Integrated by Software incorrectly |  |  |  |  |
|  | VSTDICC001; VL032018.D; Styrene; sam; 6/3/2018 2:14:26 PM; MMDadoda; 6/4/2018 10:54:54 AM; Peak Integrated by Software incorrectly. VSTDICC001; VL032018.D; Tetrachloroethene; sam; 6/3/2018 2:14:26 PM; MMDadoda; 6/4/2018 10:54:54 AM; Peak Integrated by Software incorrectly. VSTDICC0.5; VL032019.D; 1,1,1-Trichloroethane; sam; 6/3/2018 2:13:31 PM; MMDadoda; 6/4/2018 10:54:56 AM; Peak Integrated by Software incorrectly. VSTDICC0.5; VL032019.D; 1,2,4-Trichlorobenzene; sam; 6/3/2018 2:13:31 PM; MMDadoda; 6/4/2018 10:54:56 AM; Peak Integrated by Software incorrectly. VSTDICC0.5; VL032019.D; 1,2-Dichlorobenzene; sam; 6/3/2018 2:13:31 PM; MMDadoda; 6/4/2018 10:54:56 AM; Peak Integrated by Software incorrectly. VSTDICC0.5; VL032019.D; 1,4-Dioxane; sam; 6/3/2018 2:13:31 PM; MMDadoda; 6/4/2018 10:54:56 AM; Peak Integrated by Software incorrectly. VSTDICC0.5; VL032019.D; 2-Butanone; sam; 6/3/2018 2:13:31 PM; MMDadoda; 6/4/2018 10:54:56 AM; Peak Integrated by Software incorrectly. VSTDICC0.5; VL032019.D; 2-Hexanone; sam; 6/3/2018 2:13:31 PM; MMDadoda; 6/4/2018 10:54:56 AM; Peak Integrated by Software incorrectly. VSTDICC0.5; VL032019.D; Benzyl Chloride; sam; 6/3/2018 2:13:31 PM; MMDadoda; 6/4/2018 10:54:56 AM; Peak Integrated by Software incorrectly. VSTDICC0.5; VL032019.D; Chlorobenzene; sam; 6/3/2018 2:13:31 PM; MMDadoda; 6/4/2018 10:54:56 AM; Peak Integrated by Software incorrectly. VSTDICC0.5; VL032019.D; cis-1,2-Dichloroethene; sam; 6/3/2018 2:13:31 PM; MMDadoda; 6/4/2018 10:54:56 AM; Peak Integrated by Software incorrectly. VSTDICC0.5; VL032019.D; cis-1,3-Dichloropropene; sam; 6/3/2018 2:13:31 PM; MMDadoda; 6/4/2018 10:54:56 AM; Peak Integrated by Software incorrectly. |  |
|  | STDCCC010; VL032025.D; Chlorobenzene-d5; sam; 6/3/2018 2:17:15 PM; MMDadoda; 6/4/2018 3:35:44 PM; Peak Integrated by Software incorrectly. VSTDCCC010; VL032025.D; m/p-Xylene; sam; 6/3/2018 2:17:15 PM; MMDadoda; 6/4/2018 3:35:44 PM; Peak Integrated by Software incorrectly. VL0601ABL01; VL032026.D; Acetone; sam; 6/3/2018 2:17:21 PM; MMDadoda; 6/4/2018 3:35:45 PM; Peak Integrated by Software incorrectly. VL0601ABL01; VL032026.D; Methylene Chloride; sam; 6/3/2018 2:17:21 PM; MMDadoda; 6/4/2018 3:35:45 PM; Peak Integrated by Software incorrectly. VL0601ABS01; VL032027.D; Isopropyl Alcohol; sam; 6/3/2018 2:17:27 PM; MMDadoda; 6/4/2018 3:35:46 PM; Peak Integrated by Software incorrectly. VL0601ABS01; VL032027.D; m/p-Xylene; sam; 6/3/2018 2:17:27 PM; MMDadoda; 6/4/2018 3:35:46 PM; Peak Integrated by Software incorrectly. J3137-01; VL032030.D; Carbon Tetrachloride; sam; 6/3/2018 2:17:34 PM; MMDadoda; 6/4/2018 3:35:47 PM; Peak Integrated by Software incorrectly.. |  |

F. Manual Integration Comments:

Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________